MAMMOTH MTN. MET STATION

Daily Journal

*NOTE:   When finished,  save document as YYYY-MM-DD (eg,. 1999-09-17), then print document and put it in “Daily Journal” section of MMSA Met Site notebook.

| Date: | 2002-01-05 |
| --- | --- |
| Time: | 1230 |
| Name: | Kirchner |

# CONDITIONS:

| Sky: | Cloudy |
| --- | --- |
| Wind: | 7 m/s 220 deg |
| Temp: | ~ -5 C |
| Snow Depth: | I beam 303 |
| Rime? (Y or N): | N |

NOTES:

-Dug out snowboards

-Installed new phone

-Checked to see that RAS is started, affirmative

-Replaced battery and leveled PIR, checked logger it works fine

-Set shadow band declination at –22.7

-Checked CR7X1, It is working fine there is no apparent problem with the logger. Any problem is likely in the: 1.) com software/hardware (which is connected properly) or 2.) on the receiving end or 3.) problems with the phone lines (probably cross talk from wet lines).  It sounds like the other logger downloads fine which leads towards problem 1 or 2. Did this problem occur once or intermittently?

-Checked wiring on conductivity meter 8 it’s OK on logger. Perhaps there is a junction box that is wet or the wires are loose/corroded. The –99999 or –69999 readings indicate an offline or out of range (broken) sensor it intermittently records what appears to be data which could mean a connection problem. Do we have a wiring schematic for this place? If not we need one to troubleshoot this type of problem.