MAMMOTH MTN. MET STATION

Daily Journal

*NOTE:   When finished,  save document as YYYY-MM-DD (eg,. 1999-09-17), then print document and put it in “Daily Journal” section of MMSA Met Site notebook.

| Date: | 2002-12-16 |
| --- | --- |
| Time: | 11:42 |
| Name: | Bert Davis |

# CONDITIONS:

| Sky: | Overcast |
| --- | --- |
| Wind: | 9 m/s |
| Temp: | -5C |
| Snow Depth: | 1.2 |
| Rime? (Y or N): | Yes |

NOTES:  Found PC Turned off (either power out or I inadvertently turned off.)

- Snow fences buried – added new section on top of yesterday’s.

- Installed new circuit board in CR7 #1.  Jesse downloaded program and checked time.   After 1 hour, the sensors look good (e.g., Temp/RH), with exception of lower wind, and precip (of course).

- Lower wind head has missing propeller assembly.  Whole head will need replacement, but not today.

- Temp/RH sensor (old) looks like good data – tube stuffed with snow (cleared ~ 1100 AM) and sensor rimed.  Sensor rimed today as everything else – 1-2 mm clear ice.

- Precip can top blew off – replaced.  Alter shield gate blown open – replaced.  I did not replace precip can bucket.  The rime has frozen shut the precip gage door and I could not open.  I put the replacement bucket in metal box on top of the platform.

- Snow board blown away.  I found and replaced – marked it with a red bamboo stake in case it blows over again.

- Snow came into bunker through both door and Santa Claus tube.  Snow on the ozone concentration instrument.  I moved this instrument farther away from door.  This storm should take care of the problem by filling in around bunker.

- Top of Santa Claus tower lid still frozen shut.