MAMMOTH MTN. MET STATION

Daily Journal

*NOTE:   When finished,  save document as YYYY-MM-DD (eg,. 1999-09-17), then print document and put it in “Daily Journal” section of MMSA Met Site notebook.

| Date: | 11NOV2003 |
| --- | --- |
| Time: | 1430 |
| Name: | Jesse Stanley |

# CONDITIONS:

| Sky: | Clear |
| --- | --- |
| Wind: | 0 to 5 mph |
| Temp: | 1 C |
| Snow Depth: | 33 cm |
| Rime? (Y or N): | N |

NOTES:

Replaced RH sensor on Platform with existing replacement found here and it appears to be working OK.

I rejuvenated the desiccant in all radiometers, exception the one near the Belfort rain gauge. No desiccant tube existed? I couldn’t find one in any obvious places here so I may have to ship one out and have Sue install and add desiccant.

Swapped out snow depth sensor on platform I-beam as old one was reading 490, new one reads about 60. Measured ~52 cm at location very close to sensor so I adjusted the offset to 522.39. Now reads 51.94.

I checked on the Belfort rain gauge there was an inch or two of ice in the bottom and about 8 – 10 inches of snow. Cleaned out, and inspected bucket, I believe it is leaking around tube penetration through bottom of bucket, hence will no longer hold glycol. I may try to patch this if I can find something suitable.

I also made a backup directory for all the old images and daily journals at the same directory level so just the new ones will appear starting with Nov 9 2003. I did this because if the system at CRREL crashes it is annoying to have to re-upload all theses files at re-start.