MAMMOTH MTN. MET STATION

Daily Journal

*NOTE:   When finished,  save document as YYYY-MM-DD (eg,. 1999-09-17), then print document and put it in “Daily Journal” section of MMSA Met Site notebook.

| Date: | 11 Oct 2004 |
| --- | --- |
| Time: | 1330 |
| Name: | Jesse Stanley / Sharlene Stanley / Sue Burak |

# CONDITIONS:

| Sky: | Mostly Clear |
| --- | --- |
| Wind: | Gusty |
| Temp: | 6 – 7 C |
| Snow Depth: | NA |
| Rime? (Y or N): | N |

NOTES:

Overview of site. Checked bearings in upper wind speed direction sensor. Slightly noisy, but not bad. Changed as a precaution for up coming season. Will try to borrow ladder from SNARL to get to lower sensor. When Sue arrived we checked calibration of rain gauge. Seemed off. Used water after zeroing with empty bucket.  And adjusted calibration constant. Had trouble getting signal at one point. Check wiring at Campbell.  Signal tapers off at weights 11 and 12. Out to 200 + mm though.  This is less than 10 inches, les than last year. Need to see if mechanical adjustment to compensate. Wire at Cambell may have been clamped on casing a bit seemed to work afterward, and all during calibration.

Checked weights after liquid cal and recorded for future use. Left Excel spreadsheet on desktop entitled RainGauge_weights. Need to find or download from Internet, Belfort Rainguage manual.

Used vulcanizing tape to fit and seal broken PVC pipes as noted by Colee, black tape over that. Tried gluing, but didn’t seem too robust.  Mostly for water seal it appears!?.  Stood up temperature sensors that were toppled, and leveled, and weighted, should be alright to next snowfall.

Printer had issues. Checked into, but was missing black ink cartridge. Sue had brought one with her, and she changed, cleaned nozzles, etc, got it working.