MAMMOTH MTN. MET STATION

Daily Journal

*NOTE:   When finished,  save document as YYYY-MM-DD (eg,. 1999-09-17), then print document and put it in “Daily Journal” section of MMSA Met Site notebook.

| Date: | June 11, 2006 |
| --- | --- |
| Time: | 1000 |
| Name: | Sue Burak |

# CONDITIONS:

| Sky: | Clear |
| --- | --- |
| Wind: | Calm |
| Temp: | +10C |
| Snow Depth: | Wind mast 120 cm |
| Rime? (Y or N): |  |

NOTES: Ibeam sensor not reading correctly. Boom sensor is clicking, but not reading ground truth depth of 289 cm. Wind mast is 50 cm off from probe depth, adjusted offset at Campbell but sensor reads 70 cm when it shoud be 120 cm.

A couple of bare spots around north lysimeter array. Adjust shadowband.